## Slide 1
HPGe MCNP Automated Parametric Optimization Model
Air Force Institute of Technology
6 December, 2017

## Slide 2: Overview
Project Goals
Previous Work
Experiment
MCNP Model
Parametric Optimization Code
Results
Conclusions
Future Work

## Slide 3: Project Goal
Build a High-Purity Germanium (HPGe) Monte Carlo N-Particle model
Create an automated parametric optimization code to match experimental and simulated efficiencies
Perform an adjoint flux calculation using ADVANTG

## Slide 4: Previous Work
Manufacturer provided dimensions are often times not sufficient for developing models. [1]
Published relative differences between experimental and Monte Carlo simulated absolute efficiencies range from around 10% to as low as 0.2%. [2][3]
Dead layer thickness and the active volume may change over time. [4]

## Slide 5: Experiment
Performed by Lt Col O’Day (April 2017)
Eckert & Ziegler Multi-nuclide source
Count time = 24 hours
5 different source positions

## Slide 6: Detector Specifications
Canberra Standard Electrode Ge Detector (SEGe)
Model Number: GC10021
Relative Efficiency 100 %
Resolution: 1.20 keV at 122 keV and 2.1 at 1.3 MeV
Coaxial Configuration
P-Type
n+ contact
p+ contact
e-
γ

## Slide 7: Detector Specifications

## Slide 8: Experimental Setup

## Slide 9: Experimental Setup
Gamma-Ray Energy [keV] | Nuclide | Activity [µCi] | Gammas per Second
60 | Am-241 | 0.02941 | 391.7
88 | Cd-109 | 0.2707 | 363.6
122 | Co-157 | 0.01019 | 322.7
159 | Te-123 | 0.01403 | 436.1
320 | Cr-51 | 0.3389 | 1236
392 | Sn-113 | 0.05109 | 1227
514 | Sr-85 | 0.06171 | 2247
662 | Cs-137 | 0.04325 | 1362
898 | Y-88 | 0.09633 | 3347
1173 | Co-60 | 0.05101 | 1885
1333 | Co-60 | 0.05101 | 1887
1836 | Y-88 | 0.09622 | 3539
* Energies were kept constant with manufacturer provided documentation
* Source uncertainty for each energy was 3.1%

## Slide 10: Experimental Setup
(0,0)
X-axis
y-axis
A
B
C
D
E
F
G
A – 13.7385 cm
B – 3.48 cm
C – 4.8885 cm
D – 7.6 cm
E – 20.735 cm
F – 29.7385 cm
G – 9.78 cm
1
2
3
4
5

## Slide 11: MCNP Model

## Slide 12: Geometry
Origin: Centered/Bottom of shielding case
Cylinders (~14 CZs)
Planes (~20 PZs)

## Slide 13: Geometry
Top View:

## Slide 14: Materials
Material | Density [g/cm3] | Component(s)
Mylar | 1.38 | IR Window
Brass | 8.41 | Metal Clasps
Aluminum | 2.7 | Detector Housing and Casing
Germanium | 5.32 | Ge Crystal
Lithium | 0.534 | Outer Deadlayer
Boron | 2.73 | Inner Deadlayer
Copper | 8.96 | Shield Lining
Tin | 7.31 | Shield Lining
Kapton Film | 1.42 | IR Window
Air | 0.001224 | Shielding Chamber
Lead | 11.34 | Shielding
Acrylic Glass | 1.19 | Source Encapsulation
Vacuum | --- | Coaxial Space
* All materials from LANLs ACE Data Tables, or PNNLs Compendium of Material Composition Data for Radiation Transport Modeling

## Slide 15: Photon Transport
Only photons
5 MeV upper limit
No Bremsstrahlung
No Coherent scattering
No photofission
Doppler broadening
Notes: Cite manual, primer, and PNNL materials

## Slide 16: Source Definition
Isotropic point source
11 discrete energies
106 source particles
Gaussian broadening off
Tallies:
  F8 – Energy Deposition Tally
    Ge Crystal (Cell 3)
    8192 bins (10-5 - 3.14344 MeV)

## Slide 17: Parameters to Adjust
Parameter | Initial Value
Outer Top Deadlayer | 0.13 cm
Outer Sides Deadlayer | 0.13 cm
Ge Crystal Length | 8.32 cm
Kapton Window | 0.01016 cm
Inner Top Coaxial Deadlayer | 0.00003 cm
Inner Sides Coaxial Deadlayer | 0.00003 cm
Top Al Casing Thickness | 0.15 cm
Sides Al Casing Thickness | 0.15 cm
Ge Crystal Density | 5.32 g/cm3

## Slide 18: Automated Parametric Optimization Code
PICTURE
Add slides with any title/as many as you want to describe your code and the process
Include our performance metric for comparing experimental vs mcnp

## Slide 19: Results
Add 5 efficiency plots for experimental verse MCNP
  Include a table for each too
1 plot for the combined optimal parameters
Advantage results?
Include final value for relative uncertainty

## Slide 20: Conclusion
How close were we?
Does our method work?
Anything?

## Slide 21: Future Work
Reach a relative percent difference of less than 1%
  Increase adjustable parameters
  Obtain more information about internal components of HPGe
Generalize automated optimization code for other applications:
  Other detectors
  Usable on other operating systems
  Automated characterization simulations with benchmarked models

## Slide 22: Summary
Project Goals
Previous Work
Experiment
MCNP Model
Parametric Optimization Code
Results
Conclusions
Future Work

## Slide 23: Acknowledgements
Captain James Bevins, AFIT
Lt Colonel Buck O’Day, AFIT
Capt Bevins Model

## Slide 24: References
[1] R. M. Keyser, "Resolution and Sensitivity as a Function of Energy and incident Geometry for Germanium Detectors," Nuclear Instruments and Methods in Physics Research Section B: Beam Interactions with Materials and Atoms, vol. 213, pp. 236-240, 2004.
[2] R. G. Helmer, R. G. Hardy, V. E. Iacob, M. Sanchez-Vega, R. G. Neilson and J. Nelson, "The use of Monte Carlo Calculations in the Determination of a Ge Detector Efficiency Curve," Nuclear Instruments and Methods in Physics Research A, vol. 511, pp. 360-381, 2002.
[3] W. F. R. R. K. M. D. O. S. C. A. C. A. X. d. S. Guilherme J. de S. Corrêa, "COMPUTATIONAL MODELING OF A HIGH PURITY GERMANIUM," in International Nuclear Atlantic Conference, Belo Horizonte,MG, Brazil, 2011.
[4] G. F. Knoll, Radiation Detection and Measurement, Hoboken, NJ: John Wiley & Sons, Inc., 2010.
[5] D. K. P. S. J. G. M. Jeremy Lloyd Conlin, "Listing of Available ACE Data Tables," Los Alamos National Laboratory, Los Alamos National Laboratory, 2013.
[6] C. G. R. P. R. R. R. W. I. RJ McConn Jr, "Compendium of Material Composition Data for Radiation Transport Modeling," Pacific North Western National Laboratory, Pacific North Western National Laboratory, 2011.
[7] R. E. F. J. K. Shultis, "An MCNP Primer," Department of Mechanical and Nuclear Engineering, Manhattan, KS, 2011.

## Slide 25: Questions?