## Slide 1
HPGe MCNP Automated Parametric Optimization Model
Air Force Institute of Technology
6 December, 2017
Notes: HPGes/Applications
NucE world modeling for everything
Some sources are hard to predict, models provide insight to make intelligent experimental decisions, cut costs, plausibility
Raise flags on possible detector issues
Future template for future exp. With this detector

## Slide 2: Overview
Project Goals
Previous Work
Experiment
MCNP Model
Parametric Optimization Code
Results
Conclusions
Future Work
Notes: Goals, Lit, Exp, Model, optimize, results, conc.

## Slide 3: Project Goals
Design a High-Purity Germanium (HPGe) Monte Carlo N-Particle model
Create an automated parametric optimization code to match experimental and simulated efficiencies
Stretch Goal: Perform an adjoint flux calculation using ADVANTG
Notes: Model
Automated Opitmizaiton
ADVANTG adjoint Flux calc

## Slide 4: Previous Work
Manufacturer provided dimensions are often times not sufficient for developing models. [1]
Published relative differences between experimental and Monte Carlo simulated absolute efficiencies range from around 10% to as low as 0.2%. [2][3]
Dead layer thickness and the active volume may change over time. [4]
Notes: Manufacturer issues
10-0.2% range
Deadlayer/Active Volume changes

## Slide 5: Experiment
Performed by Lt Col O’Day (April 2017)
Eckert & Ziegler Multi-nuclide source
Count time = 24 hours
5 different source positions
Notes: experiment

## Slide 6: Experimental Setup
Notes: Visual, copper lining to prevent X-ray/electron backscatter

## Slide 7: Experimental Setup
Gamma-Ray Energy [keV] | Nuclide | Activity [µCi] | Gammas per Second
60 | Am-241 | 0.02941 | 391.7
88 | Cd-109 | 0.2707 | 363.6
122 | Co-157 | 0.01019 | 322.7
159 | Te-123 | 0.01403 | 436.1
320 | Cr-51 | 0.3389 | 1236
392 | Sn-113 | 0.05109 | 1227
514 | Sr-85 | 0.06171 | 2247
662 | Cs-137 | 0.04325 | 1362
898 | Y-88 | 0.09633 | 3347
1173 | Co-60 | 0.05101 | 1885
1333 | Co-60 | 0.05101 | 1887
1836 | Y-88 | 0.09622 | 3539
* Energies were kept constant with manufacturer provided documentation
* Source uncertainty for each energy was 3.1%
Notes: Sufficient Energy Range

## Slide 8: Experimental Setup
(0,0)
X-axis
Z-axis
A
B
C
D
E
F
G
A – 13.7385 cm
B – 3.48 cm
C – 4.8885 cm
D – 7.6 cm
E – 20.735 cm
F – 29.7385 cm
G – 9.78 cm
1
2
3
4
5
Notes: Visual set up

## Slide 9: Detector Specifications
Canberra Standard Electrode Ge Detector (SEGe)
Model Number: GC10021
Relative Efficiency 100 %
Resolution: 1.20 keV at 122 keV and 2.1 at 1.3 MeV
Coaxial Configuration
P-Type
n+ contact
p+ contact
e-
γ
Notes: Found on Canberra’s website

## Slide 10: Detector Specifications
Notes: Spec sheet: nominal values/polishing effects
Approximated clasp location/IR window layout
Ignored: Pin, Teflon, Power cables, top endcap clasp, rounded edges?

## Slide 11: Detector Specifications
Notes: Outer electrode – Lithium(n+)
Inner Electrode – Boron (p+)

## Slide 12: MCNP Model
Notes: **Aesthetic Effect**

## Slide 13: Geometry
Origin: Centered/Bottom of shielding case
Cylinders (~14 CZs)
Planes (~20 PZs)
Notes: Cylinders/Planes to simplify parameter/size adjustments
Interahcnageable for future models

## Slide 14: Geometry
Top View:

## Slide 15: Materials
Material | Density [g/cm3] | Component(s)
Mylar | 1.38 | IR Window
Aluminum | 2.7 | Detector Housing and Casing
Germanium | 5.32 | Ge Crystal
Lithium | 0.534 | Outer Deadlayer
Boron | 2.73 | Inner Deadlayer
Copper | 8.96 | Shield Lining
Tin | 7.31 | Shield Lining
Kapton Film | 1.42 | IR Window
Air | 0.001224 | Shielding Chamber
Lead | 11.34 | Shielding
Acrylic Glass | 1.19 | Source Encapsulation
Vacuum | --- | Coaxial Space
* All materials from LANLs ACE Data Tables, or PNNLs Compendium of Material Composition Data for Radiation Transport Modeling

## Slide 16: Parameters to Adjust
Parameter | Initial Value
Outer Top Deadlayer | 0.13 cm
Outer Sides Deadlayer | 0.13 cm
Ge Crystal Length | 8.32 cm
Kapton Window | 0.01016 cm
Inner Top Coaxial Deadlayer | 0.00003 cm
Inner Sides Coaxial Deadlayer | 0.00003 cm
Top Al Casing Thickness | 0.15 cm
Sides Al Casing Thickness | 0.15 cm
Ge Crystal Density | 5.32 g/cm3

## Slide 17: Photon Transport
Only photons
5 MeV upper limit
No Bremsstrahlung
No Coherent scattering
No photo fission
Doppler broadening
Notes: Cite manual, primer, and PNNL materials

## Slide 18: Source Definition
Isotropic point source
12 discrete energies
106 source particles
Gaussian broadening off
Tally:
  F8 – Energy Deposition Tally
    Ge Crystal (Cell 3)
    8192 bins (10-5 - 3.14344 MeV)

## Slide 19: Developer Environment
Windows 10
Required Software:
  MCNP
  Python
How-to run
Structure
  InputDeck – Location of the created MCNP Decks
  MCNP_Ouput – Location of the created data
  Model – Location of the user created MCNP deck
    Experimental Data
    FilestoMerge
  PythonCode
Notes:   InputDeck – Location of the created MCNP Decks
  MCNP_Ouput – Location of the created data
  Model – Location of the user created MCNP deck
    Experimental Data
    FilestoMerge
Required Software:
  MCNP - 6
  Python – 3
  How-to run - Known working with python installed on local C: double click the StartAutomaton.bat
  Otherwise use favorite IDE and run the python script

## Slide 20: Automated Parametric Optimization Code
Overview
Process
  Create Values for the Parameters
  Alter the MCNP Input Deck
  Run MCNP
  Compare the Results
  Record Clean and Reset
Notes: This is the code frequency, just to give some insight on how much was done. It is bloated a little bit since it is probably counting the Data Output files
Simple sweet and to the point. The created values are from 1] the iteration range and 2] the best value out of the previous parameter results, based on lowest Chi Squared value currently.

## Slide 21: Batch Files

## Slide 22: Batch Files

## Slide 23: User Changes
Notes:   Even if the user wants to simply duplicate running this file they will need to alter the runMCNP.bat with their datapath to their mcnp data and mcnp bin folders.
  First 100 lines of code involves heavy commentary and variables the user would have to alter.
  The user also has to edit the runMCNP.bat file if they change the name for the mcnpModel

## Slide 24: Parameters to Adjust
Parameter | Initial Value | Lower Bound | Upper Bound
Outer Top Deadlayer [cm] | 0.13 | 0.013 | 0.73899333
Outer Sides Deadlayer [cm] | 0.13 | 0.03 | 0.23
Ge Crystal Length [cm] | 8.32 | 7.474993997 | 9.165002
Kapton Window [cm] | 0.01016 | 0.00516 | 0.11016
Inner Top Coaxial Deadlayer [cm] | 3.00E-05 | 5.67E-06 | 0.00013
Inner Sides Coaxial Deadlayer [cm] | 3.00E-05 | 1.00E-05 | 1.00E-04
Top Al Casing Thickness [cm] | 0.15 | 0.05 | 0.25
Sides Al Casing Thickness [cm] | 0.15 | 0.05 | 0.25
Ge Crystal Density [g/cm3] | 5.32 | 5.29 | 5.32

## Slide 25: What is Created
After running MCNP the Output file is striped of its
relevant information.
Once acquired the Error and the Chi square
Can be determined
Additionally Plots are able to be created at this point

## Slide 26: Before After
Notes: Include the data from before and after showing how the default values are optimized somewhat. Also include the experimental data. Used relative error since we were trying to get near identical values compared to the given experimental data
Used an average error to simply see if any progress was made without having to look line by line.
Used a Chi Squared value ……….?

## Slide 27: Position 1
Avg Rel Error: 52.1%
Max Rel Error: 452%
Min Rel Error: 0.237%
Notes: Include all the plots explain that these are showing our results in comparison with the given experimental results. Note that the plots were also automated in their creation.

## Slide 28: Position 2
Avg Rel Error: 47.8%
Max Rel Error: 423%
Min Rel Error: 0.696%
Notes: Include all the plots explain that these are showing our results in comparison with the given experimental results. Note that the plots were also automated in their creation.

## Slide 29: Position 3
Avg Rel Error: 32.6%
Max Rel Error: 259%
Min Rel Error: 5.68%
Notes: Include all the plots explain that these are showing our results in comparison with the given experimental results. Note that the plots were also automated in their creation.

## Slide 30: Position 4
Avg Rel Error: 52%
Max Rel Error: 452%
Min Rel Error: 0.2%
Notes: Include all the plots explain that these are showing our results in comparison with the given experimental results. Note that the plots were also automated in their creation.

## Slide 31: Position 5
Avg Rel Error: 49.4%
Max Rel Error: 380%
Min Rel Error: 8.27%
Notes: Include all the plots explain that these are showing our results in comparison with the given experimental results. Note that the plots were also automated in their creation.

## Slide 32: Optimized Positions
Parameter | Initial Value | Position |  |  |  | 
 |  | 1 | 2 | 3 | 4 | 5
Outer Top Deadlayer [cm] | 0.13 |  |  |  |  | 
Outer Sides Deadlayer [cm] | 0.13 |  |  |  |  | 
Ge Crystal Length [cm] | 8.32 |  |  |  |  | 
Kapton Window [cm] | 0.01016 |  |  |  |  | 
Inner Top Coaxial Deadlayer [cm] | 0.00003 |  |  |  |  | 
Inner Sides Coaxial Deadlayer [cm] | 0.00003 |  |  |  |  | 
Top Al Casing Thickness [cm] | 0.15 |  |  |  |  | 
Sides Al Casing Thickness [cm] | 0.15 |  |  |  |  | 
Ge Crystal Density [g/cm] | 5.32 |  |  |  |  | 
Notes: Include all the plots explain that these are showing our results in comparison with the given experimental results. Note that the plots were also automated in their creation.

## Slide 33: Reasons for Possible Error
Notes: Lower energy photons are more sensitive to attenuation
Between 200-500 keV includes both the: Backscatter and annihilation peak
Backscatter – recoiled Compton scattered photon
Annihilation – 0.511 MeV photon from pair production/positron annihilation
Knolls: backscatter peak @ 184 keV for Cs-137, Comton Edge @ 662-184 = 477 keV

## Slide 34: Conclusion
The optimization method can represent the efficiency between energies of 0.392-1.836 MeV with a relative difference of X, for sources directly above the detector
Improvements are needed for-off centered sources
Unable to match low energy photon attenuation with less than a relative difference of 100%
More experimental measurements need to be taken
Further evaluation of internal components

## Slide 35: Future Work
Reach a relative percent difference of less than 1%
  Increase adjustable parameters
  Obtain more information about internal components of HPGe
  Gather more experimental data for comparison
Include more realistic radiation transport physics to reproduce experimental spectra
Apply automated optimization code to other detectors.
Explore other means for optimization

## Slide 36: Summary
Project Goals
Previous Work
Experiment
MCNP Model
Parametric Optimization Code
Results
Conclusions
Future Work

## Slide 37: Acknowledgements
Captain James Bevins, AFIT
Lt Colonel Buck O’Day, AFIT
Capt Bevins Model

## Slide 38: References
[1] R. M. Keyser, "Resolution and Sensitivity as a Function of Energy and incident Geometry for Germanium Detectors," Nuclear Instruments and Methods in Physics Research Section B: Beam Interactions with Materials and Atoms, vol. 213, pp. 236-240, 2004.
[2] R. G. Helmer, R. G. Hardy, V. E. Iacob, M. Sanchez-Vega, R. G. Neilson and J. Nelson, "The use of Monte Carlo Calculations in the Determination of a Ge Detector Efficiency Curve," Nuclear Instruments and Methods in Physics Research A, vol. 511, pp. 360-381, 2002.
[3] W. F. R. R. K. M. D. O. S. C. A. C. A. X. d. S. Guilherme J. de S. Corrêa, "COMPUTATIONAL MODELING OF A HIGH PURITY GERMANIUM," in International Nuclear Atlantic Conference, Belo Horizonte,MG, Brazil, 2011.
[4] G. F. Knoll, Radiation Detection and Measurement, Hoboken, NJ: John Wiley & Sons, Inc., 2010.
[5] D. K. P. S. J. G. M. Jeremy Lloyd Conlin, "Listing of Available ACE Data Tables," Los Alamos National Laboratory, Los Alamos National Laboratory, 2013.
[6] C. G. R. P. R. R. R. W. I. RJ McConn Jr, "Compendium of Material Composition Data for Radiation Transport Modeling," Pacific North Western National Laboratory, Pacific North Western National Laboratory, 2011.
[7] R. E. F. J. K. Shultis, "An MCNP Primer," Department of Mechanical and Nuclear Engineering, Manhattan, KS, 2011.

## Slide 39: Questions?